## techData
Data Sheet Series MKT
As of 03.06.2020
Description | MKT115-400V * | MKT115-480V-C * | MKT240-400V * | MKT240-480V-C * | MKT720-400V * | MKT720-480V-C *
Article Number | 9020-0385 | 9020-0363 | 9020-0386 | 9020-0364 | 9020-0387 | 9020-0365
Performance Data Temperature
Temperature range | -70…180 | -70…180 | -70…180 | -70…180 | -70…180 | -70…180
Temperature variation depending on setpoint [± K] | 0.2…1.8 | 0.2…1.8 | 0.1…1.0 | 0.1…1.0 | 0.3…2 | 0.3…2
Temperature fluctuation depending on setpoint [± K] | 0.1…0.6 | 0.1…0.6 | 0.1…0.4 | 0.1…0.4 | 0.1…0.5 | 0.1…0.5
Average heating-up rate according to IEC 60068-3-5 [K/min] | 5.3 | 5.3 | 5 | 5 | 4.5 | 4.5
Cooling down time from 180 °C to -70 °C [min] | 110 | 110 | 95 | 95 | 120 | 120
Average cooling down time according to IEC 60068-3-5 [K/min] | 4.2 | 4.2 | 4.2 | 4.2 | 4 | 4
Max. heat compensation at 25 °C [W] | 1500 | 1500 | 3200 | 3200 | 5000 | 5000
Electrical data
Rated Voltage [V] | 400 | 480 | 400 | 480 | 400 | 480
Power frequency [Hz] | 50 | 60 | 50 | 60 | 50 | 60
Nominal power [kW] | 5.5 | 5.5 | 6.5 | 6.5 | 13 | 13
Unit fuse [A] | 16 | 16 | 16 | 16 | 32 | 32
Phase (Nominal voltage) | 3~ | 3~ | 3~ | 3~ | 3~ | 3~
Measures
Interior volume [L] | 115 | 115 | 228 | 228 | 734 | 734
Net weight of the unit (empty) [kg] | 305 | 223 | 380 | 294 | 610 | 1
Permitted load [kg] | 60 | 60 | 70 | 70 | 160 | 160
Load per rack [kg] | 30 | 30 | 30 | 30 | 40 | 40
Viewing window width [mm] | 228 | 228 | 508 | 508 | 508 | 508
Viewing window height [mm] | 222 | 222 | 300 | 300 | 300 | 300
Wall clearance back [mm] | 300 | 300 | 300 | 300 | 300 | 300
Wall clearance sidewise [mm] | 200 | 200 | 200 | 200 | 200 | 200
Internal Dimensions
Width [mm] | 600 | 600 | 735 | 735 | 1200 | 1200
Height [mm] | 480 | 480 | 700 | 700 | 1020 | 1020
Depth [mm] | 400 | 400 | 443 | 443 | 600 | 600
Doors
Unit doors | 1 | 1 | 1 | 1 | 1 | 1
Housing dimensions not incl. fittings and connections
Width net [mm] | 980 | 980 | 1115 | 1115 | 1580 | 1580
Height net [mm] | 1725 | 1725 | 1935 | 1935 | 2005 | 2005
Depth net [mm] | 865 | 865 | 925 | 925 | 1140 | 1140
Environment-specific data
Sound-pressure level [dB(A)] | 64 | 67 | 64 | 67 | 65 | 67
Fixtures
Number of shelves (std./max.) | 1/4 | 1/4 | 1/6 | 1/6 | 1/11 | 1/11
* All technical data is specified for unloaded units with standard equipment at an ambient temperature of +22 °C ±3 °C and a power supply voltage fluctuation of ±10 %. The temperature data is determined in accordance to BINDER factory standard following DIN 12880, observing the recommended wall clearances of 10 % of the height, width, and depth of the inner chamber. Technical data refers to 100 % fan speed. All indications are average values, typical for units produced in series. We reserve the right to change technical specifications at any time.

## Description
Data Sheet Series MKT
As of 03.06.2020
Serie | MKT
Katalogbezeichnung Gruppe | Dynamic climate chambers
Series feature | Dynamic climate chambers for rapid temperature changes with extended low temperature range
Important characteristics | * Temperature range: -70 °C to 180 °C * 4 zero-voltage relay contacts * APT.line™ preheating chamber technology * Programmable condensation protection for test material * Heated viewing window with LED interior lighting * BINDER Multi Management Software APT-COM™ Basic Edition * Troubleshooting system with visual and audible alarms * Intuitive touchscreen controller with time-segment and real-time programming * Internal data logger, measured values can be read out in open format via USB * Access port with silicone plug (Models 115, 240: 50 mm, left) * 2 access ports with silicone plugs (Model 720: 80 mm, left and right) * Class 2 independent adjustable temperature safety device (DIN 12880) with visual alarm * 4 castors, two with brakes * Computer interface: Ethernet * 230 V power socket on the right-side control panel * Adjustable ramp function * Integrated chart recorder * Real-time clock * Door heating * 1 stainless steel rack * Inner chamber made of stainless steel * CFC-free refrigerants R-452A and R-23 * Cooling with cascade compressor cooling unit
Features | * Temperature range: -70 °C to 180 °C * 4 zero-voltage relay contacts * Programmable condensation protection for test material * Intuitive touchscreen controller with time-segment and real-time programming * Heated viewing window with LED interior lighting
Text of Series | Temperature ranges between -70 °C and 180 °C, with the added benefit of natural simulation, are what make the BINDER MKT series so unique. At the same time, this environmental simulation chamber fulfills the very highest precision and performance requirements for cyclical temperature tests.