## techData
Data Sheet Series UF V
As of 03.06.2020
Description | UFV500-230V * | UFV500-230V-W * | UFV500UL-240V * | UFV500UL-120V * | UFV700-230V * | UFV700-230V-W * | UFV700UL-240V * | UFV700UL-120V *
Article Number | 9020-0347 | 9020-0349 | 9020-0353 | 9020-0351 | 9020-0348 | 9020-0350 | 9020-0354 | 9020-0352
Performance Data Temperature
Temperature range | -90…-40 | -90…-40 | -90…-40 | -90…-40 | -90…-40 | -90…-40 | -90…-40 | -90…-40
Temperature variation at -80 °C [± K] | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5
Temperature fluctuation at -80 °C [± K] | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5
Cooling down time from 22 °C to -80 °C [min] | 360 | 360 | 360 | 360 | 450 | 450 | 450 | 450
Warm-up time during power failure from -80 °C / -122 °F to -60 °C / -76 °F [min] | 230 | 230 | 230 | 230 | 250 | 250 | 250 | 250
Electrical data
Rated Voltage [V] | 230 | 230 | 240 | 120 | 230 | 230 | 240 | 120
Power frequency [Hz] | 50 | 50 | 60 | 60 | 50 | 50 | 60 | 60
Nominal power [kW] | 1.6 | 1.6 | 1.6 | 1.4 | 1.6 | 1.6 | 1.6 | 1.4
Unit fuse [A] | 10 | 10 | 10 | 13 | 10 | 10 | 10 | 13
Phase (Nominal voltage) | 1~ | 1~ | 1~ | 1~ | 1~ | 1~ | 1~ | 1~
Measures
Interior volume [cu.ft.] | 16.8 | 16.8 | 16.8 | 16.8 | 24.7 | 24.7 | 24.7 | 24.7
Net weight of the unit (empty) [lbs] | 545 | 545 | 545 | 545 | 637 | 637 | 637 | 637
Permitted load [lbs] | 441 | 441 | 441 | 441 | 441 | 441 | 441 | 441
Permitted load of individual compartment [lbs] | 110 | 110 | 110 | 110 | 110 | 110 | 110 | 110
Wall clearance back [in] | 3.94 | 3.94 | 3.94 | 3.94 | 3.94 | 3.94 | 3.94 | 3.94
Wall clearance sidewise [in] | 9.85 | 9.85 | 9.85 | 9.85 | 9.85 | 9.85 | 9.85 | 9.85
Internal Dimensions
Width [in] | 23.88 | 23.88 | 23.88 | 23.88 | 35.07 | 35.07 | 35.07 | 35.07
Height [in] | 51.2 | 51.2 | 51.2 | 51.2 | 51.2 | 51.2 | 51.2 | 51.2
Depth [in] | 23.84 | 23.84 | 23.84 | 23.84 | 23.84 | 23.84 | 23.84 | 23.84
Doors
Unit doors | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Housing dimensions not incl. fittings and connections
Width net [in] | 32.54 | 32.54 | 32.54 | 32.54 | 43.7 | 43.7 | 43.7 | 43.7
Height net [in] | 77.5 | 77.5 | 77.5 | 77.5 | 77.5 | 77.5 | 77.5 | 77.5
Depth net [in] | 36.96 | 36.96 | 36.96 | 36.96 | 36.96 | 36.96 | 36.96 | 36.96
Environment-specific data
Energy consumption at set-point -80 °C and ambient temperature = 21 °C [kWh/d] | 7.9 | 7.9 | 7.9 | 7.9 | 8.1 | 8.1 | 8.1 | 8.1
Average heat dissipation at set-point -80 °C and ambient temperature = 21 °C [W] |  |  |  |  | 340 | 340 | 340 | 340
Sound-pressure level [dB(A)] | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47
Fixtures
Number of shelves (std./max.) | 3/13 | 3/13 | 3/13 | 3/13 | 3/13 | 3/13 | 3/13 | 3/13
Load-variants
Freezer racks per compartment | 4 | 4 | 4 | 4 | 6 | 6 | 6 | 6
Cryoboxes, 50 mm | 352 | 352 | 352 | 352 | 528 | 528 | 528 | 528
* All technical data is specified for unloaded units with standard equipment at an ambient temperature of +22 °C ±3 °C and a power supply voltage fluctuation of ±10%. The temperature data is determined in accordance to BINDER factory standard following DIN 12880, observing the recommended wall clearances of 10 % of the height, width, and depth of the inner chamber. All indications are average values, typical for units produced in series. We reserve the right to change technical specifications at any time.

## Description
Data Sheet Series UF V
As of 03.06.2020
Serie | UF V
Katalogbezeichnung Gruppe | Ultra low temperature freezers
Series feature | Ultra low temperature freezers with climate-neutral refrigerants
Important characteristics | * Range of temperature controller: -90 °C to -40 °C * Lowest energy consumption in its class * Powerful cascade compressor cooling unit * Environmentally friendly refrigerants R-290 and R-170 * Efficient thermal insulation with vacuum insulation panels * Innovative door gasket concept reduces buildup of ice * Ergonomic door handle * Rust-proof interior made completely of stainless steel * Detachable inner doors made of stainless steel * 3 stainless-steel shelves providing flexible placement * Water cooling (depending on version) * Zero-voltage alarm contact for in-house alarm system * Internal data logger, measured values can be read out in open format via USB * Ethernet interface * Operating hours counter * 2 access ports, Ø 28 mm, rear
Features | * Range of temperature controller: -90 °C to -40 °C * Lowest energy consumption in its class * Environmentally-friendly refrigerants R-290 and R-170 * Efficient thermal insulation with vacuum insulation panels * Rust-proof interior made completely of stainless steel
Text of Series | The BINDER ultra low temperature freezer ensures the safe storage of samples at -90 °C. It combines outstanding environmental friendliness with low energy consumption, convenient operation and an individual safety concept.